## Sheet1
Part # | Vendor | Description | QTY | Unit Price (CHF) | Price (CHF) | Assembly | Subassembly | Comment
Optical Table
RS-2000 | Newport | RS-2000 Table | 1 | 8000 | 8000 | Optical Table |  | Needs to be selected according to needs & room configuration
Detection path: Major Components
C13440-20CU V3 | Hamamatsu | ORCA-FLASH 4.0 V3 CAMLINK | 1 | 16850 | 16850 | Detection path | Camera
N1564600 | Olympus | MVX-ZB10 Zoombody | 1 | 5950 | 5950 | Detection path | Olympus Microscope
N1564200 | Olympus | MVPLAPO1x Objective | 1 | 3180 | 3180 | Detection path | Olympus Microscope | The 2x works, but has large amounts of distortion
N1563500 | Olympus | MVX-TLU Tube lens | 1 | 1170 | 1170 | Detection path | Olympus Microscope
N1564500 | Olympus | MVX-TV1xC C-Mount Adapter | 1 | 303 | 303 | Detection path | Olympus Microscope
96A350 | Ludl | 96A350 10 pos Filter Wheel | 1 | 3016.9395 | 3016.9395 | Detection path | Filter wheel
996066 | Ludl | MAC6000 Controller | 1 | 3625.3809 | 3625.3809 | Detection path | Filter wheel
73006047 | Ludl | Keypad | 0 | 461.8899 | 0 | Detection path | Filter wheel | Not strictly necessary
99A507 | Ludl | EMISSION FLANGE, OLYMPUS BX/IX Trinocular Mount | 1 | 335.5524 | 335.5524 | Detection path | Filter wheel
99A135 | Ludl | EMISSION FLANGE, OLYMPUS BX/IX Trinocular Mount | 1 | 335.5524 | 335.5524 | Detection path | Filter wheel
902-0064-000 | Nodna | Robotis Dynamixel MX-28R | 1 | 250 | 250 | Detection path | Zoom servo | Bought via nodna.de
903-0163-100 | Nodna | Robotis FR07-F101K Set Rotate Bracket (N101 Typ) | 1 | 45 | 45 | Detection path | Zoom servo | Bought via nodna.de
DX-C4-240 | Nodna | (DX-C4-240) Dynamixel DX/RX Cable 4-pin 240mm | 1 | 8 | 8 | Detection path | Zoom servo | Bought via nodna.de
P00001056 | Playzone | YF USB to RS485 Adapter | 1 | 12.5 | 12.5 | Detection path | Zoom servo | USB Adapter
Excitation path: Major Components
iChrome MLE-LFA-HP | Toptica | 405,488,561,640 nm, 100mW each, single fiber AFC | 2 | 51000 | 102000 | Excitation path | Laser | Alternatives include Omicron SOLE-6 with 2x AOM
BG-001943 | Toptica | MLE Passive Thermal Kit | 2 | 380 | 760 | Excitation path | Laser
BG-001977 | Toptica | MLE Active Thermal Kit | 0 | 1785 | 0 | Excitation path | Laser | Alternative to passive Kit for Toptica MLEs
GVS211/M | Thorlabs | 1D Large Beam (10 mm) Diameter Galvo System, Broadband Mirror (-E02), Metric, PSU Not Included | 2 | 1686.6476 | 3373.2952 | Excitation path | Scan unit | Only allows 10mm beam
GPS011 | Thorlabs | 1D or 2D Galvo System Linear Power Supply | 1 | 451.3564 | 451.3564 | Excitation path | Scan unit | Also used for the Scanlab scanners (order 2 in this case)
GCE001 | Thorlabs | Galvo Driver Card Cover | 2 | 56.1496 | 112.2992 | Excitation path | Scan unit | Only allows 10mm beam
136228 | Scanlab | Single-Axis-Module; 14/1; 450-2500nm, dynAXIS 3 M; SSV30; dynAXIS 3M without flange | 2 | 1842 | 3684 | Excitation path | Scan unit | Allows 14 mm beam, needs custom driver box, do not order for now, currently tested in Zurich
Sample & Microscope Stages
XYZRot System | Physik Instrumente | 2x M-112K033 (M-112.2DG modified for 45mm travel), 1x L-406.40DG10, 1x M-116.DG | 1 | 24800 | 24800 | Stages | Sample stage
C-884.6DC | Physik Instrumente | DC-Motor Controller, 6 CH | 1 | 5440 | 5440 | Stages | Controller
M-406.4PD | Physik Instrumente | Precision Translation Stage, 100mm | 1 | 4650 | 4650 | Stages | Detection focus
Computer
| Brentford | Intel Xeon E5-2630V3, 128 GB RAM, Geforce 960, | 1 | 7827 | 7827 | Workstation
| BenQ | BenQ BL3201PT 32" 4K LED UHD 3840x2160 | 1 | 1086 | 1086 | Screen
434494 | Digitec | Exsys EX-44094, Universal PCI-E zu 4S RS-232 Karte, (Oxford) | 1 | 86.9 | 86.9 | Microscope control |  | RS232/COM Card (Com ports are necessary for the filter wheel, possibly I stage, possibly other periphery)
6072231 | Digitec | Logitech G Saitek Farm Sim Vehicle Side Panel | 1 | 139 | 139 | Microscope control
Excitation path: Major optical components
EL-16-40-TC-VIS-5D-1-C | Optotune | EL16 Lenses | 2 | 1300 | 2600 | Excitation | ETL Unit
EL-E-4i | Optotune | Lens Driver | 2 | 270 | 540 | Excitation | ETL Unit | Have to be modified: BNC cables need to be soldered
CAB-6-300 | Optotune | Cables | 2 | 105 | 210 | Excitation | ETL Unit
| Digitec | Powered USB-Hub | 1 | 30 | 30 | Excitation | ETL Unit | Simple USB-Hub for powering the ETL-Drivers
91863 | Nikon | Nikon AF-S 50 mm f/1.4 G | 2 | 469 | 938 | Excitation | Scan unit | Have to be modified
Detection: Emission Filters
F57-405SG | AHF | QuadLine Rejectionband ZET405/488/561/640 | 1 | 761.259 | 761.259 | Detection path |  | Filters have to be selected according to lasers & needs
F39-482SG | AHF | 482/35 BrightLine HC for CFP in SG32 (405 nm excitation) | 1 | 537.7404 | 537.7404 | Detection path |  | Filters have to be selected according to lasers & needs
F37-520SG | AHF | 520/35 BrightLine HC for EGFP in SG32 (488 nm Ex) | 1 | 537.7404 | 537.7404 | Detection path |  | Filters have to be selected according to lasers & needs
F37-542SG | AHF | 542/27 BrightLine HC for YFP in SG32 | 1 | 537.7404 | 537.7404 | Detection path |  | Filters have to be selected according to lasers & needs
F76-594SG | AHF | 594 LP Edge Basic Langpass-Filter in SG32 | 1 | 610.087 | 610.087 | Detection path |  | Filters have to be selected according to lasers & needs
F76-562SG | AHF | RazorEdge Sperrfilter LP 561 RU in SG32 | 1 | 1166.184 | 1166.184 | Detection path |  | Filters have to be selected according to lasers & needs
F39-508SG | AHF | 509/22 BrightLine HC in SG32 | 1 | 620.885 | 620.885 | Detection path |  | Filters have to be selected according to lasers & needs
F37-515SG | AHF | 515/LP BrightLine HC Langpass-Filter in SG32 | 1 | 733.1842 | 733.1842 | Detection path |  | Filters have to be selected according to lasers & needs
| AHF | 535/30 ET Bandpass in SG32 | 1 | 467.7694 | 467.7694 | Detection path |  | Filters have to be selected according to lasers & needs
F47-450SG | AHF | 450/50 ET Bandpass in SG32 | 1 | 476.1918 | 476.1918 | Detection path |  | Filters have to be selected according to lasers & needs
F37-535SG | AHF | 530/43 Brightline HC in SG32 | 1 | 426.521 | 426.521 | Detection path |  | Filters have to be selected according to lasers & needs
F76-631SG | AHF | 633 LP Edge Basic in SG32 | 1 | 610.087 | 610.087 | Detection path |  | Filters have to be selected according to lasers & needs
F47-584SG | AHF | 585/40 ET Bandpass in SG32 | 1 | 476.1918 | 476.1918 | Detection path |  | Filters have to be selected according to lasers & needs
F37-565SG | AHF | 565/24 BrightLineHC in SG32 | 1 | 577.693 | 577.693 | Detection path |  | Filters have to be selected according to lasers & needs
Electronics / Signal generation
781161-01 | National Instruments | NI PXIe-1073 Integrated MXIe, 5 Periph Slots, PCIe-8361, 3m Cable | 1 | 1630 | 1630 | Signal generation
763065-01 | National Instruments | Power Cord, 220V, 10A, Swiss | 1 | 11 | 11 | Signal generation
779632-01 | National Instruments | PXI-6259 (32 Analog Inputs, 48 Digital I/O, 4 Analog Outputs) | 1 | 2800 | 2800 | Signal generation
777643-01 | National Instruments | BNC-2110 Noise Rejecting, Shielded BNC Connector Block | 3 | 520 | 1560 | Signal generation
192061-02 | National Instruments | SHC68-68-EPM Shielded Cable, 68-D-Type to 68 VHDCI Offset, 2 m | 2 | 193 | 386 | Signal generation
778512-01 | National Instruments | NI PXI-6733 with 8 16-Bit Waveform Analog Outputs | 1 | 3260 | 3260 | Signal generation
184749-02 | National Instruments | SH68-68-EP, Shielded Cable, 2 m | 1 | 186 | 186 | Signal generation
DELOCK 12498 | Reichelt | BNC to SMB cable, 3 m | 16 | 8.95 | 143.2 | Signal generation |  | SMB cables for Analog & Digital Laser control. A Sole-6 would require SMA cables
Excitation & detection path: Minor components
G022081000 | Qioptiq | Flat rail FLS 95-500-M | 2 | 119.8578 | 239.7156 | Detection Path | Detection Rail | 2nd Rail allows to park a stereomicroscope for visual inspection
G022083000 | Qioptiq | Flat rail FLS 95-1000-M | 1 | 222.4388 | 222.4388 | Detection Path | Main Rail | Main Rail the whole detection path, stages and sample rest on
G063730000 | Qioptiq | Adjustable Mirror Insert 30 | 10 | 153.3316 | 1533.316 | Excitation path | Fold mirrors | BBE1-E02 need to be glued to them
G061042000 | Qioptiq | Mounting Plate 30 | 10 | 26.995 | 269.95 | Excitation path | Fold mirrors
G061251000 | Qioptiq | Circlip Kit | 1 | 66.9476 | 66.9476 | Lab supplies | Excitation path | Special pliers in case the retaining rings of the Adjustable mirror insets 30 need to be removed
G061011000 | Qioptiq | Set of Threaded Pins M2.3x3, 150 ea. | 1 | 14.0374 | 14.0374 | Lab supplies | Excitation path
G061012000 | Qioptiq | Set of Threaded Pins M2.3x6 for Mounting Plates, 150 | 1 | 15.1172 | 15.1172 | Lab supplies | Excitation path
G061239000 | Qioptiq | Set of Special Screws M2.3x9 (20pcs.) | 1 | 30.2344 | 30.2344 | Lab supplies | Excitation path
G061238000 | Qioptiq | Set of Thumbscrews M2.3x4, 20 ea. | 1 | 24.8354 | 24.8354 | Lab supplies | Excitation path
G061111000 | Qioptiq | Set of Screws M2.3x5, (200pcs.) | 1 | 19.4364 | 19.4364 | Lab supplies | Excitation path
G061353000 | Qioptiq | Flat rail FLS 40-150 | 2 | 44.2718 | 88.5436 | Excitation path | Excitation path
G061374000 | Qioptiq | Carrier FLR 40-40 | 2 | 53.99 | 107.98 | Excitation path | Excitation relay | Mounting plate for Cube 30
G061047000 | Qioptiq | Mounting Plate 35 | 6 | 30.2344 | 181.4064 | Excitation path | Excitation relay | Mounts for 80 mm Achromats
G061081000 | Qioptiq | Cube 30 | 2 | 97.182 | 194.364 | Excitation path | Excitation relay | Mounts second to last fold mirror ahead of the scanners, have to be modified to allow a 35 mm lens to be mounted
G061225000 | Qioptiq | Holder 30 | 2 | 32.394 | 64.788 | Excitation path | Excitation relay | Mounts last fold mirror ahead of the scanners
G026529000 | Qioptiq | Carrier X 95-300, black, w/ hole pattern | 2 | 225.6782 | 451.3564 | Main Rail & Detection Path | Detection Rail | 1 for the focusing stage, 1 for the detection path
G063200000 | Qioptiq | Achr. VIS ARB2; D=31.5; F=80; mounted | 4 | 158.7306 | 634.9224 | Excitation path | Excitation relay | Relay achromats from ETL to Scanners
G061372000 | Qioptiq | Carrier FLR 40-20 | 12 | 38.8728 | 466.4736 | Excitation path | Excitation relay | 2 in the excitation path, the rest are for alignment & spares
P100/M | Thorlabs | Ø1.5" Mounting Post M6 Taps L = 100 mm | 4 | 41.78 | 167.12 | Excitation path | Excitation platform
P150/M | Thorlabs | Ø1.5" Mounting Post M6 Taps L = 150 mm | 5 | 48.1483 | 240.7414 | Excitation path | Excitation platform
TBB1515/M | Thorlabs | Large-Area Translation Stage 150 mm x 191.7 mm | 2 | 483.5992 | 967.1985 | Excitation path | Excitation platform | Platforms the whole excitation path rests on
CPVM | Thorlabs | Vertical Cage System Mounting Plate | 4 | 41.27 | 165.0798 | Excitation path | Excitation collimators & ETL
SHB025 | Thorlabs | Ø1/4" Stainless Steel Diaphragm Optical Beam Shutter with Controller | 2 | 833.8648 | 1667.7295 | Excitation path | Excitation collimators
SHCP025/M | Thorlabs | 30 mm Cage and Post Mounting Adapter for SHB025(T) Optical Beam Shutter M4 Tap | 2 | 37.3017 | 74.6034 | Excitation path | Excitation collimators | Mounting plate for shutters
SM1FCA | Thorlabs | FC/APC Fiber Adapter Plate with External SM1 (1.035"-40) Thread | 2 | 29.6297 | 59.2594 | Excitation path | Excitation collimators
CP02/M | Thorlabs | SM1-Threaded 30 mm Cage Plate 0.35" Thick 2 Retaining Rings M4 Tap | 2 | 15.7672 | 31.5345 | Excitation path | Excitation collimators | Cage plate to mount excitation collimator
ER05-P4 | Thorlabs | Cage Assembly Rod 1/2" Long Ø6 mm 4 Pack | 2 | 18.4657 | 36.9313 | Lab supplies | Excitation path
ER1-P4 | Thorlabs | Cage Assembly Rod 1" Long Ø6 mm 4 Pack | 2 | 18.4657 | 36.9313 | Lab supplies | Excitation path
ER1.5-P4 | Thorlabs | Cage Assembly Rod 1.5" Long Ø6 mm 4 Pack | 2 | 21.3122 | 42.6245 | Lab supplies | Excitation path
ER2-P4 | Thorlabs | Cage Assembly Rod 2" Long Ø6 mm 4 Pack | 4 | 22.3175 | 89.2701 | Lab supplies | Excitation path
ER3-P4 | Thorlabs | Cage Assembly Rod 3" Long Ø6 mm 4 Pack | 4 | 24.1271 | 96.5082 | Lab supplies | Excitation path
ER4-P4 | Thorlabs | Cage Assembly Rod 4" Long Ø6 mm 4 Pack | 4 | 25.9366 | 103.7463 | Lab supplies | Excitation path
SM1V10 | Thorlabs | Ø1" Adjustable Lens Tube 0.81" Travel Range | 2 | 32.0107 | 64.0213 | Excitation path | Excitation collimators
SM1L03-P5 | Thorlabs | SM1 Lens Tube 0.30" Thread Depth SM1RR Retaining Ring 5 Pack | 1 | 55.0054 | 55.0054 | Lab supplies | Excitation path
SM1L05-P5 | Thorlabs | SM1 Lens Tube 0.50" Thread Depth SM1RR Retaining Ring 5 Pack | 1 | 56.9525 | 56.9525 | Lab supplies | Excitation path
SM1L10-P5 | Thorlabs | SM1 Lens Tube 1.00" Thread Depth SM1RR Retaining Ring 5 Pack | 1 | 64.4764 | 64.4764 | Lab supplies | Excitation path
ST1XY-S | Thorlabs | XY Translator with Micrometer Drives | 2 | 355.5565 | 711.1131 | Excitation path | ETLs
XRN25C/M | Thorlabs | Compact 25 mm Travel Linear Translation Stage Rear Micrometer M6 Taps | 2 | 387.6519 | 775.3037 | Excitation path | Excitation path | Moves last fold mirror ahead of the scanners for light-sheet alignment
XRN-B1/M | Thorlabs | Baseplate for XRN25 Series Stages, Metric Slot Spacing | 2 | 30.1804 | 60.3608 | Excitation path | Excitation path
SM1A10 | Thorlabs | Adapter with External SM1 Threads and Internal C-Mount Threads | 2 | 18.4127 | 36.8255 | Lab supplies | Spare
SM1A9 | Thorlabs | Adapter with External C-Mount Threads and Internal SM1 Threads | 2 | 18.4127 | 36.8255 | Lab supplies | Spare
SM1A39 | Thorlabs | Adapter with External C-Mount Threads and External SM1 Threads | 2 | 19.6826 | 39.3652 | Excitation path | Excitation path | Connects ETLs to CPVM plate, very important
AC254-075-A-ML | Thorlabs | f=75 mm Ø1" Achromatic Doublet SM1-Threaded Mount ARC: 400-700 nm | 2 | 94.1802 | 188.3603 | Excitation path | Excitation collimators | Used as collimator to provide 15 mm beam-diameter
BBE1-E02 | Thorlabs | 1" Broadband Dielectric Elliptical Mirror 400 - 750 nm | 10 | 90.2119 | 902.1189 | Excitation path | Fold mirrors | Have to be glued to Qioptiq Adjustable Mirror Insert 30
SM2NFM | Thorlabs | Adapter with External SM2 Threads and Nikon Female F-Mount Ring | 2 | 87.5664 | 175.1328 | Excitation path | Scan unit
LCP03/M | Thorlabs | 60 mm Blank Cage Plate M4 Tap | 2 | 34.3916 | 68.7833 | Excitation path | Scan unit | Mounting plate for the 50mm excitation objective
RS2P/M | Thorlabs | Ø25.0 mm Pedestal Pillar Post M6 Taps L = 50 mm | 8 | 25.6085 | 204.8683 | Excitation path | Excitation relay and excitation platform
PS20/M | Thorlabs | Mounting Post Spacer, Height = 20 mm | 8 | 11.54 | 92.32 | Excitation path | Excitation platform
PS5M | Thorlabs | Mounting Post Spacer Height = 5 mm | 4 | 8.0953 | 32.381 | Excitation path | Excitation relay
PS3/M | Thorlabs | Mounting Post Spacer, Height = 3 mm | 8 | 11.54 | 92.32 | Excitation path | Excitation platform
RS3P/M | Thorlabs | Ø25.0 mm Pedestal Pillar Post M6 Taps L = 75 mm | 2 | 28.0424 | 56.0848 | Excitation path | Excitation relay
RS3M | Thorlabs | Ø25 mm Post Spacer Thickness = 3 mm | 6 | 7.0371 | 42.2223 | Detection path | Camera Mount
RS10M | Thorlabs | Ø25 mm Post Spacer Thickness = 10 mm | 2 | 8.1482 | 16.2963 | Lab supplies | Spare
PH50/M | Thorlabs | Ø12.7 mm Post Holder Spring-Loaded Hex-Locking Thumbscrew L=50 mm | 4 | 7.483 | 29.9321 | Lab supplies | Excitation path
PH75/M | Thorlabs | Ø12.7 mm Post Holder Spring-Loaded Hex-Locking Thumbscrew L=75 mm | 3 | 7.873 | 23.6191 | Lab supplies | Excitation path
PH150/M | Thorlabs | Ø12.7 mm Post Holder Spring-Loaded Hex-Locking Thumbscrew L=150 mm | 1 | 12.0424 | 12.0424 | Lab supplies | Excitation path
TPSM1/M | Thorlabs | Magnetic Laser Safety Screen, 200 mm x 75 mm, Metric Engraving | 6 | 32 | 192 | Laser safety | Laser safety | General-purpose magnetic beam blocks, very useful
TPSM2/M | Thorlabs | Magnetic Laser Safety Screen, 300 mm x 75 mm, Metric Engraving | 6 | 40.25 | 241.5 | Laser safety | Laser safety | General-purpose magnetic beam blocks, very useful
RA180/M-P5 | Thorlabs | Right-Angle End Clamp for Ø1/2" Posts, M6 Stud and 5 mm Hex, 5 Pack | 2 | 57.5 | 115 | Lab supplies | Excitation path + Spare
RA90/M-P5 | Thorlabs | Right-Angle Clamp for Ø1/2" Posts, 5 mm Hex, 5 Pack | 2 | 51.01 | 102.02 | Lab supplies | Excitation path + Spare
TR50/M-P5 | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 50 mm, 5 Pack | 4 | 24.41 | 97.64 | Lab supplies | Excitation path + Spare
TR75/M-P5 | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 75 mm, 5 Pack | 2 | 26.16 | 52.32 | Lab supplies | Excitation path + Spare
TR100/M-P5 | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 100 mm, 5 Pack | 2 | 27.62 | 55.24 | Lab supplies | Excitation path + Spare
TR150/M-P5 | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 150 mm, 5 Pack | 2 | 33.06 | 66.12 | Lab supplies | Excitation path + Spare
TR250/M | Thorlabs | Ø12.7 mm Optical Post SS M4 Setscrew M6 Tap L = 250 mm | 1 | 8.5926 | 8.5926 | Detection Path | Servo mount
PH50/M-P5 | Thorlabs | Ø12.7 mm Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=50 mm, 5 Pack | 1 | 40.25 | 40.25 | Lab supplies | Excitation path + Spare
PH75/M-P5 | Thorlabs | Ø12.7 mm Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=75 mm, 5 Pack | 1 | 43.23 | 43.23 | Lab supplies | Excitation path + Spare
PH100/M-P5 | Thorlabs | Ø12.7 mm Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=100 mm, 5 Pack | 1 | 47.93 | 47.93 | Lab supplies | Excitation path + Spare
BA2/M-P5 | Thorlabs | Mounting Base, 50 mm x 75 mm x 10 mm, 5 Pack | 1 | 34.34 | 34.34 | Lab supplies | Excitation path + Spare
BA1S/M-P5 | Thorlabs | Mounting Base, 25 mm x 58 mm x 10 mm, 5 Pack | 1 | 24.46 | 24.46 | Lab supplies | Excitation path + Spare
PSHA/M | Thorlabs | Adjustable Height Collar M6 Locking Screw | 1 | 61.6404 | 61.6404 | Detection path | Camera Mount
C1515/M | Thorlabs | Ø1.5" Mounting Post Bracket M6 Taps | 1 | 95.2384 | 95.2384 | Detection path | Camera Mount
XT95P3 | Thorlabs | Base Plate for 95 mm Rails | 1 | 48.6774 | 48.6774 | Detection path | Camera Mount | Needs to be modified
XT95P11/M | Thorlabs | Drop-On Rail Carriage for 95 mm Rails, 95.0 mm Long, M6 Tapped Holes | 1 | 80.23 | 80.23 | Sample mounts | Baseplate for Birdhouse
KB25/M | Thorlabs | Complete 25 mm x 25 mm Kinematic Base Top and Bottom Plates M4 Counterbores | 10 | 74.3388 | 743.3883 | Sample mounts | Sample holder | It's always good to have more of these - depending on the number of samples & cuvettes - also for intermediate storage of samples
CP20S | Thorlabs | 30 mm Cage System Iris Ø20.0 mm Maximum Aperture | 2 | 83.5981 | 167.1962 | Excitation path | Excitation collimators
CF125C/M-P5 | Thorlabs | Clamping Fork 31.5 mm Counterbored Slot M6 x 1.0 Captive Screw 5 Pack | 1 | 49.1324 | 49.1324 | Lab supplies | Excitation path
PH20/M | Thorlabs | 12.7 mm Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=20 mm | 2 | 6.8351 | 13.6703 | Sample holder | Sample holder
TR20/M-P5 | Thorlabs | Stainless Steel Optical Posts - Metric | 1 | 20.7322 | 20.7322 | Sample holder | Sample holder
M-360-90 | Newport | Angle Bracket, 90°, Slotted Faces, Metric | 1 | 82 | 82 | Detection path | Camera bracket
M-CXL95-120 | Newport | Rail Carriage, 120 mm Length, M5 and M6 Thread, X95 Series | 4 | 127 | 508 | Sample mounts | Sample holder and general supplies | Sample holder and general supplies - it is useful to have more of these for exchanging the outer cuvettes
M-CXL95-50 | Newport | Rail Carriage, 50 mm Length, M6 Thread, X95 Series | 2 | 95 | 190 | Detection path | Servo holder & spare | Servo holder & spare
RD-PDT-B | Radiant Dyes | Tip-Tilt Plate | 2 | 280 | 560 | Sample holder | Sample holder | Order one for each planned immersion cuvette mount (e.g. 40x40 or 50x50)
Alignment Tools
SI254 | Thorlabs | Shearing Interferometer with a 10-25.4 mm Beam Diameter Shear Plate | 1 | 488 | 488
SI100P | Thorlabs | Shear Plate, 5-10 mm Beam Diameter | 1 | 184 | 184
SICP | Thorlabs | 30 mm Cage System Adapter for Shearing Interferometers with Diffuser Viewing Screen | 1 | 76 | 76
CCM1-G01/M | Thorlabs | 30 mm Cage Cube-Mounted Protected Aluminum Turning Mirror, M4 Tap | 1 | 184 | 184
SM1FC | Thorlabs | FC/PC Fiber Adapter Plate with External SM1 (1.035"-40) Thread | 1 | 31 | 31
SM1FCA | Thorlabs | FC/APC Fiber Adapter Plate with External SM1 (1.035"-40) Thread | 1 | 32.78 | 32.78
AC254-100-A-ML | Thorlabs | f=100 mm Ø1" Achromatic Doublet SM1-Threaded Mount ARC: 400-700 nm | 1 | 94.1802 | 94.1802
SM1V15 | Thorlabs | Ø1" Adjustable Lens Tube, 1.31" Travel Range | 1 | 36.8 | 36.8
SM1M20 | Thorlabs | SM1 Lens Tube Without External Threads, 2" Long, Two Retaining Rings Included | 1 | 17 | 17
SM1T2 | Thorlabs | SM1 (1.035"-40) Coupler, External Threads, 0.5" Long | 2 | 21.21 | 42.42
DGM05 | Thorlabs | Digital Indicator | 1 | 304.75 | 304.75 |  |  | Very useful
DIGC6 | Thorlabs | Digital Calipers | 1 | 148.21 | 148.21 |  |  | Very useful
R2L2S3P2 | Thorlabs | Grid Distortion Target, 1.5" x 1.5", 250 µm Grid Spacing | 1 | 348.45 | 348.45 |  |  | Very useful
FFM1 | Thorlabs | 30-mm-Cage-Compatible Rectangular Filter Mount | 1 | 60.95 | 60.95
DG10-1500-H1-MD | Thorlabs | Ø1" SM1-Mounted Frosted Glass Alignment Disk w/Ø1 mm Hole | 1 | 34.5 | 34.5 |  |  | Nice to have
DG20-1500-H2-MD | Thorlabs | Ø2" SM2-Mounted Frosted Glass Alignment Disk w/Ø2 mm Hole | 1 | 51.45 | 51.45 |  |  | Nice to have
CPA1 | Thorlabs | 30 mm Cage System Alignment Plate with Ø1 mm Hole | 2 | 13.8 | 27.6
LCPA1 | Thorlabs | 60 mm Cage Alignment Plate | 1 | 19.09 | 19.09
LCP01/M | Thorlabs | 60 mm Cage Plate, SM2 Threads, 0.5" Thick, M4 Tap (Two SM2RR Retaining Rings Included) | 2 | 40.51 | 81.02 |  |  | Nice to have as spares
CP20S | Thorlabs | 30 mm Cage System Iris Ø20.0 mm Maximum Aperture | 2 | 83.5981 | 167.1962 |  |  | Nice to have
SM1D12D | Thorlabs | SM1 Ring-Actuated Iris Diaphragm (Ø0.8 - Ø12 mm) | 4 | 70.104 | 280.416 |  |  | Nice to have
G061008000 | Qioptiq | Mounting Plate Without Hole | 2 | 28.75 | 57.5 |  |  | Have to be modified with SM1 threads
G061372000 | Qioptiq | Carrier FLR 40-20 | 2 | 38.8728 | 77.7456
| Fisso | Strato M-28 movable arm | 1 | 205 | 205 |  |  | Very useful
| Fisso | Thread adapter M8/M6 thread reduction for Fisso | 2 | 12 | 24 |  |  | Very useful
General lab supplies
LAB-16 | Newport | Ultra Jet Canned Air Duster, 10 oz. (300ml) | 4 | 16.5 | 66 |  |  | Or other canned duster
Optical Wonder | Baader Planetarium | Microfiber cloth, Optical Wonder | 4 | 7.47 | 29.88 |  |  | General cleaning cloth for non-risky surfaces (Eyepieces etc)
19569/16OZ | Cargille | Immersion oil from cargille for Clarity: Code 50350 nD=1.4587 fused silica matching oil | 2 | 140 | 280 |  |  | A decent supply is recommend (though it is not consumed much)
13023MPRM2.5 | Kummer | EPO-TEK 302-3M in 2.5 g pre-pack, burst seal, 25x packs (minimum order) | 1 | 438 | 438 |  |  | 1 2.5g package is sufficent to glue 10 mirrors, Fabian will take care of that
MC-50E | Thorlabs | Lens Tissues, 25 Sheets per Booklet, 50 Booklets in a Closeable Box | 1 | 85.1 | 85.1
ESK11 | Thorlabs | BNC Adapters Essentials Kit, 62 Pieces | 1 | 914 | 914 |  |  | This or any other BNC-Kit with all the common connectors (e.g. T, L, Straight etc.)
CA2848 | Thorlabs | SMA Coaxial Cable, SMA Male to BNC Male, 48" (1219 mm) | 2 | 18 | 36 |  |  | Sometimes useful
2249-C-120 | Thorlabs | 0.3 m BNC Cable | 4 | 15.25 | 61
2249-C-24 | Thorlabs | 0.6 m BNC Cable | 10 | 16.5 | 165
2249-C-60 | Thorlabs | 1.5 m BNC cable | 10 | 24 | 240
2249-C-120 | Thorlabs | 3m BNC cable | 6 | 24 | 144
BKF12 | Thorlabs | Black Aluminium Foil | 2 | 32 | 64 |  |  | Sometimes useful
PLT2 | Thorlabs | Ø1/2" (Ø12.7 mm) Photoluminescent Stickers, Qty: 100 | 1 | 25 | 25 |  |  | Sometimes useful for marking setups in darkness
T743-2.0 | Thorlabs | High-Performance Black Masking Tape, 2" x 180' (50 mm x 55 m) Roll | 2 | 51.58 | 103.16
PSS7 | Thorlabs | 7-Piece Precision Screwdriver Set | 1 | 60 | 60 |  |  | Or other, cheaper sets
HW-KIT6/M | Thorlabs | Imperial Screw Thread Adapter Kit | 1 | 136.48 | 136.48
HW-KIT6/M | Thorlabs | Metric Screw Thread Adapter Kit | 1 | 136.48 | 136.48
HW-KIT1/M | Thorlabs | M4 Cap Screw and Hardware Kit | 1 | 60.95 | 60.95
HW-KIT3/M | Thorlabs | M4 Setscrew and Hardware Kit | 1 | 119.416 | 119.416
HW-KIT4/M | Thorlabs | M6 Setscrew and Hardware Kit | 1 | 105.8 | 105.8
HW-KIT2/M | Thorlabs | M6 Cap Screw and Hardware Kit | 2 | 119.6 | 239.2
TC3/M | Thorlabs | 15-Piece Balldriver and Hex Key Kit with Stand, Metric | 2 | 92 | 184
TC2 | Thorlabs | 20-Piece Balldriver and Hex Key Kit with Stand, Imperial | 1 | 96.23 | 96.23
SPW801 | Thorlabs | Adjustable Spanner Wrench | 1 | 102.44 | 102.44
SPW602 | Thorlabs | SM1 Spanner Wrench, Graduated, Length = 3.88" | 2 | 28 | 56
|  | Multiport-power outlets | 4
|  | Swiss power cords to standard cold-device plugs | 10
|  | Cable wraps | 1 |  |  |  |  | Lots of them
|  | Isopropanol p.a. for lens cleaning | 1
Custom components
| Mechanical Workshop | Galvomount-22mm-diameter-Thorlabs-Left.ipt | 1 |  |  | Excitation Path |  | Only if Thorlabs Scanners are used | Aluminium
| Mechanical Workshop | Galvomount-22mm-diameter-Thorlabs-Right.ipt | 1 |  |  | Excitation Path |  | Only if Thorlabs Scanners are used | Aluminium
| Mechanical Workshop | Galvomount-22mm-diameter-Left-Scanlab-V4.ipt | 0 |  |  | Excitation Path |  | Only if Scanlab Scanners are used, parts under test at HIFO, will require additional parts | Aluminium
| Mechanical Workshop | Galvomount-22mm-diameter-Right-Scanlab-V4.ipt | 0 |  |  | Excitation Path |  | Only if Scanlab Scanners are used, parts under test at HIFO, will require additional parts | Aluminium
| Mechanical Workshop | 60mm-blank-cageplate-LCP03M-Nikon-MOD.ipt | 2 |  |  | Excitation Path |  | Modified LCP03M | Aluminium
| Mechanical Workshop | MVX-10-Bracket-V5.ipt | 1 |  |  | Detection Path |  |  | Aluminium
| Mechanical Workshop | Servo-Dynamixel-Adapter-V2.ipt | 1 |  |  | Detection Path |  |  | Brass
| Mechanical Workshop | Large-cuvette-mount-30mm-V1.ipt | 0 |  |  | Sample Mounts |  | if 30 mm cuvettes are desired | POM (Polyoxymethylen)
| Mechanical Workshop | Large-cuvette-mount-40mm-V3.ipt | 2 |  |  | Sample Mounts |  | Order one for each planned immersion cuvette mount / imaging medium | POM (Polyoxymethylen)
| Mechanical Workshop | Large-cuvette-mount-50mm-V3.ipt | 2 |  |  | Sample Mounts |  | Order one for each planned immersion cuvette mount / imaging medium | POM (Polyoxymethylen)
| Mechanical Workshop | M-116-Adapter.ipt | 1 |  |  | Sample Mounts |  |  | Aluminium
Will be redesigned | Mechanical Workshop | Sample-cuvette-clamp-20mm-V3.ipt | 10 |  |  | Sample Mounts |  | Redesign is being worked on | POM (Polyoxymethylen)
| Mechanical Workshop | Sample-Stage-Gate | 1 |  |  | Sample Stage Gate |  |  | Aluminium
| Mechanical Workshop | Cube_30_with_35mm_hole | 2 |  |  | Excitation Path |  |  | Aluminium
| Mechanical Workshop | M605-to-X95-adapter-V1 | 1 |  |  | Detection Path |  |  | Aluminium
| Mechanical Workshop | X95carrier-to-M605-adapter-V1 | 1 |  |  | Detection Path |  |  | Aluminium
| Mechanical Workshop | XT95P3_X95_Base_plate_mod | 1 |  |  | Detection Path |  | Modified XT95P3 | Aluminium
Sample cuvettes
| Portmann Instruments | UQ-205-H80 10x20x80 | 0 | 518.5 | 0 |  |  | Whole-Mouse CNS Cuvettes
| Portmann Instruments | UQ-753-H100 40x40x100 | 0 | 764 | 0
| Portmann Instruments | UQ-205-H120 10x20x120 | 0 | 816 | 0 |  |  | Whole-Mouse CNS Cuvettes
| Portmann Instruments | UQ-753-H120 40x40x120 | 0 | 952 | 0
| Portmann Instruments | UQ-751 20x20 mm | 0 | 99.5 | 0
| Portmann Instruments | UQ-752 30x30 | 0 | 128.5 | 0
| Portmann Instruments | UQ-753 40x40 | 0 | 156 | 0
| Portmann Instruments | UQ-754 50x50 | 0 | 184.5 | 0
| Portmann Instruments | UQ-203 5x10x45 | 0 | 106.5 | 0
| Portmann Instruments | UQ-204 10x10x45 | 6 | 106.5 | 639
| Portmann Instruments | UQ-205 10x20x45 | 10 | 121.5 | 1215 |  |  | Standard cuvette for mouse brains & CLARITY
Fun Stuff
| sygns.de | Custom Neon ¦ mesospim.org ¦ Classic Red ¦ 100cm on white matte Aluminium | 1 | 932 | 932 | Neon Sign |  | This is the glowing red "mesospim.org" neon sign, can be identical to order DE176398
|  |  |  | SUM | 246115.865